## Start here
Medical Device Threat Model Template
Blue Goat Cyber. Free to use. Fill in each tab in order. This template follows how we build and review threat models; the FDA does not require this exact format. Send the finished file for a gap assessment at bluegoatcyber.com/threat-model-gap-assessment.
1. Scope | Describe the device version you plan to submit and the assumptions you rely on.
2. Harm list | One shared list of patient and system harms, taken from your ISO 14971 file. Both the safety and security sides use these IDs.
3. Acceptance criteria | Write down which risk levels are acceptable BEFORE you rate any threat.
4. Elements & entry points | List every data flow diagram element and every interface an attacker could reach. Mark trust boundaries.
5. Threat table | One row per threat. Fill all eight columns. Status is a tracking flag only.
6. Initial risk guide | How to rate column 7: exploitability combined with severity of harm, rated as if no control exists yet.
7. Mitigation strategies | How to write column 8: "Strategy: the control." Eliminate, Mitigate, Accept or Transfer.
Example | A finished four-row table for a fictional device. Copy the shape, not the content.
What this is not | The threat table gives the initial, pre-control rating. Detailed scoring, residual risk and the controlled/uncontrolled decision belong in your cybersecurity risk assessment, joined back by Threat ID.

## 1 Scope
Scope and assumptions
Describe the device as it will ship, not an early design.
Device name and model
Software/firmware version covered
Date and revision of this threat model
Intended use and use environment (hospital, home, clinic)
Device class and submission pathway
Hardware components
Software components (apps, cloud, backend)
Update and patch path
Third-party software (link to SBOM)
Architecture views included (global system, multi-patient harm, updatability/patchability, security use cases)
Assumptions (network, physical access, user roles)
Out of scope, and why

## 2 Harm list
Shared harm list
Copy harms from your ISO 14971 file. Severity uses your existing safety scale; do not create a separate security severity scale for patient harm.
Harm ID | Harm | Severity (your safety scale) | Hazard ID(s) in ISO 14971 file
HARM-01 | Example: overdose / hypoglycemia | 5 Catastrophic | HAZ-012
HARM-02 | Example: delayed or wrong diagnosis from falsified data | 4 Critical | HAZ-020

## 3 Acceptance criteria
Acceptance criteria (write before scoring)
Decide and sign off what each initial risk level requires before you rate any threat. Replace the examples with your own approved criteria.
Initial risk | What it requires | Approved by / date
Very High | Example: not acceptable. Layered controls required; known exploited vulnerabilities must be fixed or removed, never accepted.
High | Example: control required and verified by a test case.
Medium | Example: control required unless a documented justification is approved.
Low | Example: may be accepted with a recorded, owned decision.

## 4 Elements & entry points
Elements and entry points
Every element of your data flow diagram and every interface an attacker could reach. Name the actual surface: "USB service port", not "the network".
Entry ID | Element type | Name | Entry point / interface | Crosses trust boundary?
Entry-1 | Data flow | Device-to-cloud upload | TLS upload to cloud API | Yes: device to cloud

## 5 Threat table
Threat table: the eight minimum columns
One row per distinct threat. Never reuse or renumber Threat IDs. Column 5 must end in a patient or system harm. Every row needs a test case. Column 7 is the INITIAL rating, as if the control were not there.
1 Threat ID | 2 Element | 3 Entry point | 4 STRIDE | 5 Threat and result of exploitation | 6 Analysis / test case | 7 Threat risk (initial) | 8 Mitigation(s): "Strategy: control" | Status
THREAT-001
THREAT-002
THREAT-003
THREAT-004
THREAT-005
THREAT-006
THREAT-007
THREAT-008
THREAT-009
THREAT-010
THREAT-011
THREAT-012
THREAT-013
THREAT-014
THREAT-015
THREAT-016
THREAT-017
THREAT-018
THREAT-019
THREAT-020
THREAT-021
THREAT-022
THREAT-023
THREAT-024
THREAT-025
THREAT-026
THREAT-027
THREAT-028
THREAT-029
THREAT-030
THREAT-031
THREAT-032
THREAT-033
THREAT-034
THREAT-035
THREAT-036
THREAT-037
THREAT-038
THREAT-039
THREAT-040
THREAT-041
THREAT-042
THREAT-043
THREAT-044
THREAT-045
THREAT-046
THREAT-047
THREAT-048
THREAT-049
THREAT-050
THREAT-051
THREAT-052
THREAT-053
THREAT-054
THREAT-055
THREAT-056
THREAT-057
THREAT-058
THREAT-059
THREAT-060
THREAT-061
THREAT-062
THREAT-063
THREAT-064
THREAT-065
THREAT-066
THREAT-067
THREAT-068
THREAT-069
THREAT-070
THREAT-071
THREAT-072
THREAT-073
THREAT-074
THREAT-075
THREAT-076
THREAT-077
THREAT-078
THREAT-079
THREAT-080
THREAT-081
THREAT-082
THREAT-083
THREAT-084
THREAT-085
THREAT-086
THREAT-087
THREAT-088
THREAT-089
THREAT-090
THREAT-091
THREAT-092
THREAT-093
THREAT-094
THREAT-095
THREAT-096
THREAT-097
THREAT-098
THREAT-099
THREAT-100

## 6 Initial risk guide
How to rate column 7 (initial risk)
Exploitability stands in for probability, because a deliberate attacker has no failure rate. Rate exploitability from the five CVSS v4.0 exploitability metrics (attack vector, attack complexity, attack requirements, privileges required, user interaction). Severity is the harm named in column 5 on your ISO 14971 scale. Patient harm is the primary axis: a catastrophic harm rates high even at modest exploitability. The grid below is an example; use the one in your approved acceptance criteria.
Severity \ Exploitability | Low | Medium | High
5 Catastrophic | High | Very High | Very High
4 Critical | Medium | High | Very High
3 Serious | Low | Medium | High
2 Minor | Low | Low | Medium
1 Negligible | Low | Low | Low
Exploitability rule of thumb: High if a working exploit is circulating (for example a CISA KEV listing) or every metric is at its easiest value. Low if the attack is not reachable over a network, or needs both high privileges and high complexity. Otherwise Medium. Try it: bluegoatcyber.com/tools/cvss-v4-calculator

## 7 Mitigation strategies
Writing column 8
Write each mitigation as "Strategy: the control." It is the plan, written before it is proven. It is not the residual risk, the test result, or the signed acceptance record. Higher harm means more independent layers.
Strategy | When to use it | Example
Eliminate | The feature is not essential. Remove the surface entirely. Strongest option. | Eliminate: disable USB mass storage and autorun; remove the debug port from production units
Mitigate | The feature must stay. Add a control that makes it hard to abuse. Most common. | Mitigate: TLS 1.2+ with certificate pinning; sign the payload; reject on signature failure
Accept | Residual is already low, or a control costs more than it helps. Always a recorded, owned decision with sign-off. | Accept: verbose non-sensitive logging on a physically secured port; signed off by the risk owner
Transfer | The control genuinely lives with someone else. Record the assumption you now depend on. | Transfer: rely on hospital network segmentation (documented assumption), plus device-side TLS

## Example
Example: fictional imaging kiosk
A finished table for a fictional device (Northwind CardioScan). Copy the shape, not the content.
1 Threat ID | 2 Element | 3 Entry point | 4 STRIDE | 5 Threat and result of exploitation | 6 Analysis / test case | 7 Threat risk (initial) | 8 Mitigation(s): "Strategy: control" | Status
THREAT-001 | Process | Kiosk touchscreen | T, E | Operator escapes the kiosk shell to the OS through an unrestricted dialog. Shell access on the acquisition host, and scans can be altered at the source. | KIOSK-001 | High | Mitigate: harden the kiosk allow-list, disable OS shortcuts and dialogs. | Open
THREAT-002 | Process | USB port | T, I | Malicious USB device injects input or mounts storage to overwrite or copy scans. The saved scan is overwritten before upload, and PHI is copied off the device. | USB-001 | High | Mitigate: disable mass storage and HID autorun, port allow-list. Transfer: rely on hospital physical controls for the service port. | Open
THREAT-003 | Data store | Local scan store | T | Scan file edited on disk before upload, with no integrity check. A falsified scan is uploaded and trusted by the cloud. | FS-001 | High | Mitigate: sign the scan payload at capture, verify the signature on ingest. | Open
THREAT-004 | Data flow | Upload flow (TLS) | S, T | TLS downgraded or intercepted on the hospital network. The scan is rewritten in transit and the clinician acts on a falsified scan. | NET-001 | High | Mitigate: enforce TLS 1.2 or higher with certificate pinning; reject downgrade. | Open